## גיליון1
Thermocouples
Tolerances on Initial Values of EMF vs. Temperature
Reference Junction 32ºF (0ºC)
Tolerances (whcihever is greater)
Special |  | standard |  | Temperature Range |  | Calibration Type
ºC | ºF | ºC | ºF | ºC | ºF
Thermocouples
(±0.25%) | (2) | (±0.5%) | (2) | (870 to 1700) | 1600 to 3100 | B
(±1.0 or ±0.4%) | (2) | (±1.7 or ±0.5%) | (2) | (0 to 870) | 32 to 1600 | E
(±1.1 or ±0.4%) | (2) | (±2.2 or ±0.75%) | (2) | (0 to 760) | 32 to 1400 | J
(±1.1 or ±0.4%) | (2) | (±2.2 or ±0.75%) | (2) | (0 to 1260) | 32 to 2300 | K or N
(±0.6 or ±0.1%) | (2) | (±1.5 or ±0.25%) | (2) | (0 to 1480) | 32 to 2700 | R or S
(±0.5 or ±0.4%) | (2) | (±1.0 or ±0.75%) | (2) | (0 to 370) | 32 to 700 | T
(5) | (2) | (±1.7 or ±1%) | (2) | (-200 to 0) | -328 to 32 | E
(5) | (2) | (±2.2 or ±2%) | (2) | (-200 to 0) | -328 to 32 | K
(5) | (2) | (±1.0 or ±1.5%) | (2) | (-200 to 0) | -328 to 32 | T
Extension Wires
(±1.0) | ±1.8 | (±1.7) | ±3.0 | (0 to 200) | 32 to 400 | EX
(±1.1) | ±2.0 | (±2.2) | ±4.0 | (0 to 200) | 32 to 400 | JX
(±1.1) | ±2.0 | (±2.2) | ±4.0 | (0 to 200) | 32 to 400 | KX or NX
(±0.5) | ±0.9 | (±1.0) | ±1.8 | (0 to 100) | 32 to 200 | TX
Compensating Extension Wires
* | * | (±4.2) | ±7.6 | (0 to 200) | 32 to 400 | BX
* | * | (±6.8) | ±12.2 | (0 to 260) | 32 to 500 | CX
* | * | (±5.0) | ±9.0 | (0 to 200) | 32 to 400 | RX, SX
(1) Tolerances in this table apply to new essentially homogeneous thermocouple wire, normally in the size range 0.25 to 3 mm in diameter (No.30 to No.8 AWG) and used at temperatures not exceeding the recommended limits on page 20.If used at higher temperatures these tolerances may not apply.
(2). At a given temperature that is expressed in ºC , the tolerance expressed in ºF is 1.8 times larger than the tolerance expressed in ºC. Note: Wherever applicable, percentage-based tolerances must be computed from temperatures that are expressed in ºC.
(3) Caution: Users should be aware that certain characteristics of thermocouple materials, including the EMF vs. temperature relationship may change with time in use, consequently, test results and performance obtained at time of manufacture may not necessarily apply throughout an extended period of use. Tolerances given above apply only to new wire as delivered to the user and do not allow for changes in charactedstics with use. The magnitude of such changes will depend on such factors as wire size, temperature, time of exposure end environment, It shoold be further noted that due to possible changes in homogeneity, attempting to recalibrate used thermocouples is likety to yield irrelevant resutts, and is not recommended. However, it may be appropriate to compare used thermocouples in-situ with new or known good ones to ascertain their suitability for further service under he conditions of the comparison.
(4) Thermocouples and thermocouple materials are normally supplied to meet the tolerances specified in the table for temperatures above 0ºC. The same materials, however. may not fall within the tolerances given for temperatures below ºC in the second section of the table. If materials are required to meet the tolerances stated for temperatures below ºC the purchase order must so state. Selection of materials usually will be required.
(5). Special tolerances for temperatures below ºC are difficult to justify due to limited available information. However, the following values for Types E and the T thermocouples are suggested as a guide for discussion between purchaser and supplier: Type E: -200 to 0ºC ±1.0ºC or ±0.5% (whichever is greater): Type T: -200 to 0ºC ±0.5 or ±0.8% (vvtiiohever is greater). Initial values of tolerance tar Type J thermocouples at temperatures below 0ºC and special toterances for Type K Thermocouples below 0ºC are not given due to the characteristics of the materials.
(6) Tolerances in the table represent the maximum error contribution allowable from new and essentially homogeneous thermocouple extension wire when exposed to the full temperature range given above. Extension grade materials are not intended for use outside the temperature range shown
(7) Thermocouple extension wire makes a contribution to the total thermoelectric signal that is dependent upon the temperature difference between the extreme ends of to extension wire length. The actual magnitude of any error introduced into a measuring circuit by homogeneous and correctly connected extension wires is equal to the algebraic difference of the deviations at its two end temperatures, as determined for that extension wire pair.
(8)Tolerances in the table apply to new and essentially homogeneous thermocouple compensating extension wire when used at temperatures within the range given above.
(9) Thermocouple ocmpensating extension wire makes a contribution to the total thermoelectric signal that is dependent upon the temperature difference between the extreme ends of the compensating extension wire length.
(10) Special compensating extension wires are not necessary with Type B over the limited temperature range 32 to 125ºF (0 to 50ºC), where the use of non-compensated (copper/copper) conductors introduces no signilicant error. For a somewhat larger temperature gradient of 32 to 210ºF (0 to 100ºC) across the extension portion of the circuit, the use of non-compensated (copper/copper) extension wires may result in small errors, the magnitude of which will not exceed the tolerance values given in the table above for measurements above 1800ºF (1000ºC). Proprietary alloy compensating extension wire is available for use over 32 to 400ºF (0 to 2OOºC) temperature range.
* Special tolerance grade compensating extension wires are not available